Example - SCHEME OF WORK WITH PROCEDURE

School of Physics, University of Exeter

Laboratory/Room/Activity

These guidelines have been prepared by..................................................

Before commencing work with Class 3B/4 lasers, you must read this document, and sign the sheet at the end to confirm this, and your agreement to abide by the protocols contained herein.

Purpose and Structure of this document

The principal aim of this document is to outline the elements of good laser practice as they apply specifically to experiments currently being undertaken in the above laboratory. General aspects of laser safety are covered in sections of the manuals accompanying the lasers.

The document is structured as follows: at the top level (this sheet) an overall description is made of laser research activity in this laboratory. The user is then referred to a number of accompanying documents under two headings: Laser types in use and Safety Protocols.

Description of Activity:

.......Description of the activity/experiment and its purpose..............................

Four types of situation have been identified which require separate safety protocols, where appropriate:

- Setting up;

- Adding new elements;

- Day-to-day operation of the experiment; and

- Maintenance/servicing of the laser system.

Separate sections dealing with each of these are to be included where appropriate.

Lasers in use

......... Reference to completed Hazard/Risk Assessment proforma ...................................

As a general rule all Class 3B/4 laser emissions are capable of causing severe eye damage if viewed directly, or as a specular (i.e., mirror-like) reflection. Control measures (careful planning, beam pipes, blocking of reflections, safety eyewear) must be taken to avoid this!

Authorised users of the above lasers are (to be named):

Supervisor(s):

Student(s):

Postdoctoral Researcher(s):

Others:

1.	Laser safety protocol: Setting up

Definition

Setting up applies to the initial installation of a new experiment, and to major changes such as the addition of a new type of laser system, or, for example, a complete change of beam paths.

Protocol/Scheme of Work

Planning:

- The installation or changes should be discussed with a supervisor prior to operation of Class 3B/4 laser systems.

- In the case of a completely new experiment, the School Laser Safety Officer and/or the University Laser Safety Adviser must be consulted and invited to visit the lab.

- The laser beam paths and associated optics should be planned to minimise the possibility of stray reflections.

- Termination of each main laser beam should be planned.

- Provision of suitable laser safety eyewear should be addressed at this stage.

Initial safety checks:

- Before starting the Class 3B/4 lasers, beam paths should be inspected for any objects that should not be there.

- Laser warning signs should be activated, where installed and unauthorised persons excluded and doors closed.

- Alignment may be carried out by one or at most two authorised laser operators. No one else may be present in the room during this procedure and watches, bracelets and other reflective jewellery should be removed.

- Appropriate laser safety eyewear should be worn if practicable. Visible or multi-wavelength alignment may have to be carried out without laser safety eyewear, as it would otherwise be impossible to visualise the laser beam on a card. In this case, extra caution must be exercised by the operator(s).

Initial alignment and suppression of stray reflections:

- Initial laser beam alignment should be performed with a Class 2 or Class 1 alignment laser (e.g., He-Ne or small CW diode laser). Remember that the final beam path may differ slightly due to dispersion (i.e., the beam path may be slightly wavelength-dependent)

- At this stage, each and every optic element in the beam path must be analysed for stray reflections. Initially this can be done by predicting the likely path of specular (i.e., non-diffuse) reflections and the actual reflections of the Class 2/1 alignment laser may also be used to help identify stray reflections.

- Suitable beam blocks, opaque at the appropriate wavelengths, should then be installed when necessary to block all stray reflections.

- ‘Beam pipes’ should be installed at this stage to cover longer runs of laser beam, and especially any beams that leave the confines of the laser table. It is recognised that there may be some places where beam pipes are inappropriate, e.g. when the distance between optics is very short. Beam pipes should be designed to allow limited access to the beam for alignment checking without removal.

Alignment using Class 3B/4 lasers at low power

- The next stage of alignment using the Class 3B/4 lasers may be carried out only after obtaining the verbal permission of a supervisor.

- Alignment may be carried out by one or at most two authorised laser operators. No one else may be present in the room during this procedure and watches, bracelets and other reflective jewellery should be removed.

- Under no circumstances must direct viewing of the laser beam be attempted even if the beam has been attenuated. The use of a video camera for remote viewing should be considered.

- All optics should be checked for damage, and the stability of optics mounts verified prior to operation of laser.

- This next stage in alignment should be carried out using the lowest possible laser energy (e.g. operating a Nd:YAG laser on fixed-Q) at which it is possible to visualise the laser beam in an appropriate fashion. The method of visualisation is dependent on the wavelength: for UV or visible light, the beam can be viewed on a fluorescent card. An invisible infrared beam may be visualised using LCD heat sensitive paper or possibly using burn paper or a laser power meter.

- In the case of UV or IR beams, appropriate laser safety eyewear should be worn during the alignment procedure at all times when the laser pulse energy exceeds the MPE: note that it should not block the wavelength-shifted visible fluorescence (UV) or the heat effect on LCD paper or burn paper (IR), which can then be used to visualise the beam.

- Visible or multi-wavelength alignment may have to be carried out without laser safety eyewear, as it would otherwise be impossible to visualise the laser beam on a card. In this case extra caution must be exercised by the operator(s) after appropriate consideration of alternatives and assessment of risks. The laser(s) should be operated below the MPE if possible and in any case at the lowest practicable pulse energy. Blocking of possible stray reflections must be double-checked prior to carrying out this stage.

- Alignment of each laser beam to variable diameter apertures (iris diaphragms) should be employed where possible to minimise the necessity for multi-wavelength alignment.

- Further alignment at full power may be carried out in accordance with the protocol outlined under ‘day-to-day operation’.

2.	Laser safety protocol: Adding new elements

Definition

Adding new elements applies to the introduction of any new optic such as a lens or filter into the beam path of a Class 3B/4 laser.

Protocol/Scheme of work

Planning:

- The placement of additional optics should be planned to minimise the possibility of stray reflections.

- Beam blocks should be devised to terminate any unavoidable stray reflections.

Initial safety checks:

- Before starting the Class 3B/4 lasers, beam paths should be inspected for any objects that should not be there and beam pipes should be replaced if necessary.

- Laser warning signs should be activated, unauthorised persons excluded and laboratory doors closed.

- Alignment may be carried out by one or at most two authorised laser operators. No one else may be present in the room during this procedure and watches, bracelets and other reflective jewellery should be removed.

- Appropriate laser safety eyewear should be worn if practicable. Visible or multi-wavelength alignment may have to be carried out without laser safety eyewear, as it may otherwise be impossible to visualise the laser beam on a card. In this case extra caution must be exercised by the operator(s).

- All optics should be checked for damage, and stability of optics mounts verified.

Initial alignment and suppression of stray reflections:

- Once a new optic is in place, initial alignment should be performed with a Class 1/2 alignment laser (e.g., He-Ne or small CW diode laser). For simple optics it may be judged sufficient to proceed to the next step without using a Class 1/2 alignment laser.

- The new optic element in the beam path must be analysed for stray reflections. This can be done by predicting the likely path of specular (i.e., non-diffuse) reflections. The actual reflections of the Class 1/2 alignment laser may also be used to help identify stray reflections.

- Suitable beam blocks, opaque at the appropriate wavelengths, must then be installed to block all these stray reflections.

- Any effect ‘downstream’ of the new optic should be checked. ‘Beam pipes’ should be re-installed at this stage.

Alignment using Class 3B/4 lasers at low power:

- This may now be carried out in accordance with the procedure outlined under ‘setting up’ with the exception that explicit permission of a supervisor is not deemed necessary for addition of a simple optical element. (Anything more complex should be taken as ‘setting up’ and the protocol followed accordingly.)

3.	Laser safety protocol: Day-to-day running

Definition

Day-to-day running applies to the operation of Class 3B/4 lasers under all circumstances except setting up or addition of a new optic element. It includes initial, minor realignment of laser beams at the beginning of an experimental run and ‘tweaking’ of alignments during an actual experiment.

Protocol/Scheme of work

Initial safety checks:

- Before starting the Class 3B/4 lasers, beam paths should be inspected for any objects that should not be there, and beam pipes should be replaced if necessary.

- Laser warning signs should be activated, unauthorised persons excluded, and laboratory doors closed.

- Alignment may be carried out by one or at the most two authorised laser operators. No one else may be present in the room during this procedure and watches, bracelets and other reflective jewellery should be removed.

- Appropriate laser safety eyewear should be worn if practicable. Visible or multi-wavelength alignment may have to be carried out without laser safety eyewear, as it would otherwise be impossible to visualise the laser beam on a card. In this case extra caution must be exercised by the operator(s).

- All optics should be checked for damage, and stability of optics mounts verified.

Check using Class 3B/4 lasers at low power:

- Before turning on full power, the beam path of each laser should be verified in turn, using the lowest possible pulse energy and visualising the beam in an appropriate fashion (e.g., on fluorescent card or using video camera).

- Under no circumstances must direct viewing of the laser beam be attempted even if the beam has been attenuated. There must be no exceptions to this rule.

Minor realignment (‘tweaking’) with lasers running at full power:

- During an experimental run, it will sometimes be necessary to re-optimise the alignment to recover lost signal. Of necessity, this can only be carried out at full power, with all lasers on. Extra caution should therefore be exercised.

- All beam guards/pipes and blocks for stray reflections should remain in place during this procedure. Beam pipes should be designed to allow limited access to the beam for alignment checking without removal.

- It is especially important to wear appropriate laser safety eyewear when visualising laser beams at full power. As before, in a visible or multi-wavelength experiment this may not be practicable, and extra caution should therefore be exercised.

- It may be possible (and indeed, preferable) to apply minor ‘tweaks’ to the alignment using the experimental signal as a guide. In this case it is not necessary to visualise the laser beams.

- If it is necessary to visualise the laser beams, e.g. to optimise beam overlap, this may be done by an appropriate safe method e.g. by fluorescent card or by remote viewing using a video camera. If practicable the output power should be restricted e.g. by the use of iris diaphragms.

4.	Laser safety protocol: Maintenance

Definition

Maintenance applies to adjustments or procedures specified in the user information provided by the manufacturer/supplier with the laser product, which are performed by the user for the purposes of assuring the intended performance of the product. It may require access within the laser enclosure. Any entry into the laser system enclosure potentially exposes the laser worker to additional non-optical hazards, for instance those associated with high voltages, and toxic chemicals, in addition to accessing high energy laser beams that are normally enclosed. The activities can be very diverse, and specific protocols will need to be prepared for each laser, this document provides some general guidelines on the planning of Maintenance.

NB Manufacturer’s type servicing is outside the scope of this document

Protocol/Scheme of work

Planning:

- Before commencing the maintenance, the manual for the laser system should be consulted, to identify the recommended procedure.

- n the case of anything other than routine maintenance, and/or when the laser manual does not give a procedure, the advice of a laser technician should be sought. Some procedures should only be conducted by an experienced laser technician.

- The risks associated with the procedure should be assessed, the control measures reviewed, and the conclusions recorded. In the case of some regular maintenance procedures (such as changing the dye solution in a dye laser), reference to an existing protocol may well suffice.

- Maintenance involving the alignment of a laser beam inside a laser enclosure, for instance introducing the pump laser into a dye laser, can lead to an increased laser radiation exposure risk, since part of the beam path of a normally enclosed, and potentially very high power beam is likely to be open. The protocol for Setting Up should be consulted.

A SIGNATURE SHEET WILL ACT AS A LIST OF AUTHORISED USERS AND A RECORD THAT PERSONS HAVE BEEN GIVEN INFORMATION AND AGREE TO FOLLOW THE PROCEDURES LAID OUT.

Signature Sheet

This sheet must be signed by all workers before commencing work with Class 3B/4 lasers.

By signing below, you confirm that you have read this document, and you agree to abide by the protocols and guidelines contained herein.